# GW Instek GSP-9300頻譜分析儀 ASK解調及分析功能DEMO

這次的演示是GSP-9300 頻譜分析儀的ASK解調及分析功能。

這 DEMO KIT是個ASK及FSK的信號產生器。此DEMO KIT有3種預設頻率，3種波特率及可設定輸出為連續或叢發模式。

架設方式：

- 連接DEMO KIT到GSP-9300前面板的USB接口

- 用SMA RF CABLE連接DEMO KIT輸出端到GSP-9300的輸入端。

儀器操作:

- 首先設定DEMO KIT的參數.

- 按Option Control鍵 > Demo Kit[F4]鍵>Demo Kit[F1]鍵*2次,啟動DEMO KIT.

- 按Modulation [F2]鍵, 選擇ASK[F1]鍵

- 按Frequency[F3]鍵,選擇433MHz[F2]鍵

- 按Baudrate[F5]鍵, 選擇波特率為4800[F1]鍵

- 按Output Mode[F6]鍵, 選擇10101010字串,Seq5[F7].

- DEMO KIT的參數設定結束.

- 按 Frequency 鍵> Center Freq[F1]鍵>4,3,2,點,9,9,2,MHz[F2]鍵, 設定中心頻率為432.992MHz

- 按 Span鍵> Span[F1]鍵 >5,0,0,kHz[F3]鍵, 設定頻段為500kHz

- 在頻譜分析儀上可以觀測到ASK信號.

- 開始設定ASK解調及分析功能觀測ASK的信號.

- 按 Measure鍵> Demod[F2]鍵>ASK[F4]鍵>ASK Analysis[F1]*2次,啟動ASK分析功能

- ASK分析測量參數，包括調幅深度，調變速率，載波功率，載波頻率偏移，信納比，符號碼(symbol)及波形(waveform)。用戶可以設定調幅深度，載波功率和載波偏移的合格/失敗測試結果。

- 按Symbol Rate[F2]鍵，輸入4,8,0,0,s [F2]，設定符號碼率為4800

- 目前螢幕上半部的測試結果為ASK解調的波形(waveform).用戶可以按View[F4]鍵，選擇符號碼(symbol)[F1].

- 用戶可以在螢幕上看到ASK解調的符號碼為之前設定的Seq5,10101010字串.

- 按Return[F7]鍵 > 按More　1/2[F7]鍵 > 按Limit[F4]鍵

- 按AM Depth[F2],設定9,5,Enter[F1]

- 按Carrier offset[F3],設定１,kHz[F3]

- 按Limit[F1]*2次，啟動合格/失敗結果測試 > 按Return[F7]*2次

- 按View[F4]鍵，選擇波形(waveform)[F2] > 按Return[F7] > 按More 1/2[F7]

- 按AF Trigger[F1]>Trigger mode[F3]*2次,設定單次觸發>Trigger Level[F4],設定觸發準位為9,0,Enter[F1] (%)

- 按Action Now[F6], 啟動單次觸發。GSP-9300暫停測量ASK信號。

- 以上是GSP-9300 ASK分析功能的介紹.

- 感謝您的收看

# GW Instek GSP-9300 Spectrum Analyzer ASK Demodulation and Analysis Function DEMO

Now, I am going to demonstrate the ASK demodulation and analysis function of GSP-9300 spectrum analyzer.

This demo kit is an ASK and FSK signal generator. In this demo kit, there are three preset frequency, three Baud rates and selectable continuous or burst output mode.

First of all, connect demo kit with the USB port located at the front panel of GSP-9300. Next, connect demo kit output terminal with GSP-9300 input terminal with SMA RF cable.

- To begin the operation, first, input demo kit parameters.

- Press option control key, demo kit (F4) key, demo kit (F1) key twice to activate demo kit

- Press modulation (F2) key and select ASK (F1) key

- Press frequency (F3) key and select 433 MHz (F2) key

- Press Baud rate (F5) key and select baud rate of 4800 (F1) key

- Press output mode (F6) key and select binary code 10101010,and Seq5(F7)

- This is the end of demo kit parameters input. Next, set up the ASK demodulation & analysis function to observe ASK signals.

- Press frequency key, center freq (F1) key, input 4, 3, 2, dot, 9, 9, 2 MHz (F2) key to set center frequency as 432.992 kHz.

- Press span key, span (F1) key, input 5, 0, 0, kHz (F3) key to set frequency bandwidth as 500kHz.

- Now, ASK signals are shown on the screen.

- Next, set ASK demodulation and analysis function to observe ASK signals.

- Press measure key, Demod (F2) key, ASK Analysis (F4) , ASK Analysis (F1)twice to activate ASK analysis function

- ASK Signal Analysis measures parameters including frequency deviation, modulation rate, carrier power, carrier frequency offset, Signal-to-noise and distortion ratio (SINAD), symbol and waveform. Users can set the criterion in frequency deviation, carrier power and carrier offset for pass/fail test result determination.

- Press symbol rate(F2) key, input 4,8,0,0 s (F2) to set symbol rate as 4800

- Now, the top half of the screen shows the test result of waveform demodulated by ASK. Users press View (F4) and select symbol (F1)

- Now, the previous set ASK demodulation symbol of binary code Seq5,10101010 is shown on the screen

- Press Return (F7), and More 1/2(F7), press Limit (F4).

- Press AM Depth (F2), input 9, 5, Enter (F1).

- Press Carrier offset(F3),input1,kHz(F3)

- Press Limit(F1) twice to activate pass/fail test result, press Return(F7) twice

- Press View(F4), select waveform(F2), press Return(F7), and More 1/2 (F7)

- Press AF Trigger (F1), Trigger mode (F3) twice, Trigger level (F4), input 9,0 Enter (F1) to set AF Trigger and Trigger Level as 90%

- Press Action Now (F6) to activate AF Trigger. GSP-9300 stops measuring ASK signals

- This concludes the demonstration of GSP-9300 ASK analysis function.

- Thank you very much for your interest in GSP-9300.