固緯電子 GSP-9300 VS 是徳科技 N9320B頻譜分析儀的掃瞄速度比較

這次的演示是固緯電子GSP-9300與是徳科技 N9320B頻譜分析儀的掃瞄速度比較。

架設方式：

- 用BNC 測試線及轉接頭連接GSP-9300及N9320B前面板的射頻輸入端.

- 連接DEMO KIT到GSP-9300前面板的USB接口

- 用SMA 射頻測試線及BNC對SMA轉接頭連接DEMO KIT輸出端到GSP-9300及N9320B的BNC測試線。

儀器操作:

- 首先設定DEMO KIT的參數.

- 按Option Control鍵 > Demo Kit[F4]鍵>Demo Kit[F1]鍵*2次,啟動DEMO KIT. 此DEMO KIT預設為FSK信號.

- DEMO KIT的參數設定結束.

- 先開始設定GSP-9300頻譜分析儀的參數.

- 按 Frequency 鍵> Center Freq[F1]鍵>3鍵>1鍵>5鍵> MHz[F2]鍵, 設定中心頻率為315MHz

- 按 Span鍵> Span[F1]鍵 >1鍵> MHz[F2]鍵, 設定頻段為1MHz

- 在頻譜分析儀上可以觀測到2FSK信號.

- GSP-9300目前的掃瞄模式為正常模式 (normal mode), 掃瞄時間為64ms.

- 用戶可以按Sweep Mode鍵,切換至快速模式(fast mode),加快掃瞄時間,掃瞄時間變更為983us.

- 用戶可以用此快速掃瞄模式捕獲瞬態的FSK信號,以呈現正確的信號波形.

- 接著設定N9320B頻譜分析儀的參數.

- 按 Frequency 鍵>選Center Freq鍵>3鍵>1鍵>5鍵> 選MHz鍵, 設定中心頻率為315MHz

- 按 Span鍵> >1鍵> 選MHz鍵, 設定頻段為1MHz

- 在頻譜分析儀上可以觀測到2FSK信號.

- N9320B的掃瞄時間為21.5ms.

- 用戶在相同的設定參數下, N9320B無法捕獲瞬態的FSK信號及無法呈現正確的信號波形.

- 以上是是固緯電子GSP-9300與是徳科技 N9320B頻譜分析儀的特性比較.

- 感謝您的收看

GW Instek GSP-9300 & Keysight N9320B Spectrum Analyzer Sweep Speed Comparison

Now, I am going to demonstrate the sweep speed comparison between GW Instek GSP-9300 and Keysight N9320B Spectrum Analyzer.

First of all, connect RF input terminals of GSP-9300 and N9320B with BNC cable and adaptors. Connect DEMO Kit to the USB port located at the front panel of GSP-9300. Next, connect demo kit output terminal and the BNC cable of GSP-9300 and N9320B with SMA RF cable and BNC to SMA adaptor.

- To begin the operation, first, input demo kit parameters.

- Press option control key, demo kit (F4) key, demo kit (F1) key twice to activate demo kit. This DEMO KIT, by default, is FSK signals.

- This is the end of demo kit parameters input.

- Next, input GSP-9300 parameters.

- Press frequency key, center freq (F1) key, input 3, 1, 5, MHz (F2) key to set center frequency as 315 MHz.

- Press span key, span (F1) key, input 1 MHz (F2) key to set frequency bandwidth as 1 MHz.

- Now, 2FSK signals are shown on the screen.

- The sweep mode of GSP-9300 is currently on normal mode, sweep time is 64ms.

- Users can press Sweep Mode to switch to fast mode so as to expedite sweep time to 983us.

- Users can obtain transient FSK signals by using fast mode to show correct signal waveform.

- Next, input N9320B spectrum analyzer parameters.

- Press frequency key, center freq key, input 3, 1, 5, MHz to set center frequency as 315 MHz

- Press span key, input 1 MHz to set frequency bandwidth as 1 MHz

- Now, 2FSK signals are shown on the screen

- The sweep time of N9320B is 21.5ms

- Under the same parameters, N9320B can not obtain transient FSK signals, therefore, correct signal waveform can not be revealed.

- This concludes the sweep speed comparison between GW Instek GSP-9300 and Keysight N9320B Spectrum Analyzer.

- Thank you very much for watching.