# GW Instek GSP-9300頻譜分析儀 FM分析功能DEMO

這次的演示是GSP-9300 頻譜分析儀的FM分析功能。

這 DEMO KIT是個ASK及FSK的信號產生器。此DEMO KIT有3種預設頻率，3種波特率及可設定輸出為連續或叢發模式。

架設方式：

- 連接DEMO KIT到GSP-9300前面板的USB接口

- 用SMA RF CABLE連接DEMO KIT輸出端到GSP-9300的輸入端。

儀器操作:

- 首先設定DEMO KIT的參數.

- 按Option Control鍵 > Demo Kit[F4]鍵>Demo Kit[F1]鍵*2次,啟動DEMO KIT.

- 按Modulation [F2]鍵, 選擇FSK[F2]鍵

- 按Frequency[F3]鍵,選擇433MHz[F2]鍵

- 按Deviation[F4]鍵,選擇50kHz[F2]鍵

- 按Baudrate[F5]鍵, 選擇波特率為9600[F2]鍵

- 按Output Mode[F6]鍵, 選擇 normal mode.

- DEMO KIT的參數設定結束.開始設定FM分析功能觀測FSK的信號.

- 按 Frequency 鍵> Center Freq[F1]鍵>4鍵>3鍵>3鍵> MHz[F2]鍵, 設定中心頻率為433MHz

- 按 Span鍵> Span[F1]鍵 >5鍵>0鍵>0鍵> kHz[F3]鍵, 設定頻段為500kHz

- 在頻譜分析儀上可以觀測到FSK信號.

- 按 Measure鍵> Demod[F2]鍵>FMAnalysis [F2]>FMAnalysis[F1]*2次,啟動FM分析功能

- FM分析測量參數，包括頻偏，調製速率，載波功率，載波頻率偏移和信納比。用戶可以設定頻偏，載波功率和載波偏移的判斷條件來進行成功/失敗測試結果。

- 按LimitEdit[F5]鍵> Freq.Deviation [F1]鍵>1鍵>0鍵>0鍵>kHz[F3]鍵,設定頻偏為100kHz.

- 按Return[F7]鍵

- 按Pass/Fail[F6]*2次,可在螢幕上看到Pass/Fail訊息.

- 按Setup[F2]>IF Bandwidth[F1]>1鍵>0鍵>0鍵>kHz[F3], 設定IF Bandwidth中頻頻寬為100kHz.

- 按Time Axis[F3]>Scale/Div[F3]>2鍵>0鍵>0鍵>us[F4], 設定時域軸為每格200us

- 按Return[F7]

- 按Deviation Axis[F4]>Scale/Div[F3]>2鍵>0鍵>kHz[F3],設定頻偏軸為每格20kHz

- 按Return[F7]

- 按AF Trigger[F5]>Trigger Setup[F1]>Trigger mode*2次,選擇單次觸發>Trigger Level[F3]>2鍵>0鍵>kHz[F3],設定單次觸發及觸發準位為20kHz

- 按Ｒeturn[F7]

- 按Action Now[F5], 啟動單次觸發。GSP-9300暫停測量FSK信號。

- 按Stop Time[F4], 逆時針轉動滾輪scroll,可以達到放大zoom in的效果。

- 以上是GSP-9300 FM分析功能的介紹.

- 感謝您的收看

# GW Instek GSP-9300 Spectrum Analyzer FM Analysis Function DEMO

Now, I am going to demonstrate the FM analysis function of GSP-9300 spectrum analyzer.

This demo kit is an ASK and FSK signal generator. In this demo kit, there are three preset frequency, three Baud rates and selectable continuous or burst output mode.

First of all, connect demo kit with the USB port located at the front panel of GSP-9300. Next, connect demo kit output terminal with GSP-9300 input terminal with SMA RF cable.

- To begin the operation, first, input demo kit parameters.

- Press option control key, demo kit (F4) key, demo kit (F1) key twice to activate demo kit

- Press modulation (F2) key and select FSK (F2) key

- Press frequency (F3) key and select 433 MHz (F2) key

- Press deviation (F4) key and select frequency deviation of 50kHz (F2) key

- Press Baud rate (F5) key and select baud rate of 9600 (F2) key

- Press output mode (F6) key and select normal mode

- This is the end of demo kit parameters input. Next, set up the FM analysis function to observe FSK signals.

- Press frequency key, center freq (F1) key, input 4, 3, 3, MHz (F2) key to set center frequency as 433 MHz.

- Press span key, span (F1) key, input 5, 0, 0, kHz (F3) key to set frequency bandwidth as 500kHz.

- Now, 2FSK signals are shown on the screen.

- Press measure key, Demod (F2) key, FM Analysis (F2) , FM Analysis (F1)twice to activate FM analysis function.

- FM Signal Analysis measures parameters including frequency deviation, modulation rate, carrier power, carrier frequency offset and Signal-to-noise and distortion ratio (SINAD). Users can set the criterion in frequency deviation, carrier power and carrier offset for pass/fail test result determination.

- Press limitedit (F5) key, freq deviation (F1) key, input 5,0, kHz (F3) key to set FSK frequency deviation as 50kHz

- Press return (F7) key

- Press pass/fail (F6) twice. Pass/Fail indicator is shown on the screen

- Press Setup (F2), IF Bandwidth (F1), input 1,0,0 kHz (F3) to set IF Bandwidth as 100kHz.

- Press Time Axis (F3), Scale/Div (F3), input 2,0,0 us (F4) to set Time Axis as 200 us per frame.

- Press return (F7) key

- Press Deviation Axis (F4), Scale/Div (F3), input 2,0, kHz (F3) to set Deviation Axis as 20 kHz.

- Press return (F7) key

- Press AF Trigger (F5), Trigger Setup (F1), Trigger mode twice, select single trigger, Trigger level (F3), input 2,0 Enter (F1) to set AF Trigger and Trigger Level as 20kHz

- Press return (F7) key

- Press Action Now (F5) to activate AF Trigger. GSP-9300 stops measuring FSK signals.

- Press Stop Time (F4), turn Scroll counter clockwise to achieve zoom in effect.

- This concludes the demonstration of GSP-9300 FM analysis function.

- Thank you very much for your interest in GSP-9300.