# GW Instek GSP-9300頻譜分析儀功能示範—測量10MHz參考輸出信號

# GSP-9300頻譜分析儀的後面板有一個10MHz的參考頻率輸出信號源.用戶可以透過BNC測試線及轉接頭,連接參考頻率(Ref out)輸出端至前面板的射頻輸入(RF Input)端.

# 操作的步驟如下:

# 按Preset鍵,將GSP-9300頻譜分析儀設定在預設值(default)設置上

# 按Frequency鍵>Center Freq[F1]鍵,輸入1,0,MHz[F2]鍵.設定中心頻率為10MHz.

# 按Span鍵>Span[F1]鍵,輸入1,5,MHz[F2]鍵. 設定頻寬為15MHz.

# 按Amplitude鍵>Ref Level[F1]鍵,輸入1,0,dBm[F1]鍵. 設定參考準位為10dBm.

# 按Peak Search鍵, 就可以快速搜尋信號的峰値點.

# 在螢幕上可見標記marker顯示10MHz頻率及功率值.

# 接著啟動及設定不同的軌跡以記錄軌跡的最大及最小値.

# 按Trace鍵>select Trace[F1]>Trace2[F2]>MaxHold[F3], 啟動軌跡2,設定為最大保持以記錄信號每次出現的最大値.

# 按select Trace[F1]>Trace3[F3]>MinHold[F4], 啟動軌跡3,設定為最小保持以記錄信號每次出現的最小値.

# 以上就是頻譜分析儀測量10MHz參考輸出信號的介紹.謝謝您的收看.

# I am going to demonstrate GW Instek GSP-9300 Spectrum Analyzer Function Demonstration-How to measure 10MHz Reference Output signals?

# On the rear panel of the GSP-9300 spectrum analyzer, there is a 10MHz reference frequency output signal source. Users can use BNC cable and adaptor to connect Ref out with RF Input on the front panel.

# To begin the demonstration:

# Press Preset, set GSP-9300 spectrum analyzer on default

# Press Frequency, Center Freq[F1],input 1,0,MHz[F2] to set center frequency as 10MHz.

# Press Span, Span[F1],input 1,5,MHz[F2] to set bandwidth as 15MHz.

# Press Amplitude, Ref Level[F1],input 1,0,dBm[F1] to set reference level as 10dBm.

# Press Peak Search to rapidly search signal’s  peak value.

# Now, marker is shown on the screen displaying 10MHz frequency and power value

# Next, activate and set different traces to record the maximum and minimum value of trace.

# Press Trace, select Trace[F1], Trace2[F2], MaxHold[F3] to activate trace 2 and set maximum hold to record the maximum value for each signal.

# Press select Trace[F1], Trace3[F3], MinHold[F4], to activate trace 3 and set minimum hold to record the minimum value for each signal.

# This concludes 10 MHz reference output signal measurement demonstration by using a GSP-9300 spectrum analyzer.Thank you very much for watching.